## Slide 1
Time-Domain Systems and THz Material Characterization
EE4730 Oral Exam | Time-Domain Topic | Daniel Tyukov | July 2026
Notes: Good morning, I'm Daniel Tyukov. Topic: time-domain, the photoconductive spectrometer, and identifying six unknown dielectric slabs.

## Slide 2: Why measure in the time domain
• The THz band (0.3 to 3 THz) sits above transistor fmax (records around 0.75 THz) and below optical sources. Generating and detecting power is the hard part.
• Two generation paradigms from the lectures: multiply a microwave oscillator up in frequency, or switch a DC bias with a femtosecond laser in a photoconductor.
• The photoconductive route gives a single picosecond pulse whose spectrum covers the entire band at once.
• Measure the field E(t), take an FFT, and one scan characterizes a material from 0.15 to 3 THz. That is what this lab does.
 | Heterodyne | Photoconductive
Source | LO chain, xN | fs laser + DC bias
Signal | CW tone | ps pulse
Bandwidth | ~30 GHz | up to ~2 THz
Use here | VNA labs | this lab
Lecture 2: the two ways to make THz signals
TDS spectrometer chain (lab manual)
Notes: THz sits above transistors, below optical sources. Two routes: multiply a microwave oscillator up, clean tone but ~30 GHz bandwidth; or DC-bias a femtosecond laser, one-ps pulse, whole band. One scan, one FFT: 0.15-3 THz.

## Slide 3: The photoconductive antenna
• An LT-GaAs gap between two electrodes. The 1.59 eV band gap matches the 780 nm pump laser.
• The femtosecond pulse creates carriers, the DC bias accelerates them, and recombination cuts the current off after about 1 ps. That current burst radiates the THz pulse.
• The lecture photocurrent model has exactly three ingredients: carrier generation (Gaussian laser pulse), Drude acceleration, and exponential recombination with τc.
• The receiver is the same device without bias. The incident THz field accelerates the carriers, so the gated current is proportional to the field at the moment of illumination.
• At high optical power the radiated field screens the bias and the photocurrent saturates. That is why high-power sources use arrays of matched PCAs.
PCA model and Norton equivalent (Lecture 2)
Notes: PCA: LT-GaAs gap between two electrodes. 1.59 eV gap, 780 nm pump. Laser creates carriers, bias accelerates, recombination kills current after ~1 ps, that burst radiates the pulse. Three ingredients: generation, acceleration, recombination. Receiver: same device, no bias; current proportional to field at that instant. High power: field screens bias, saturates, hence PCA arrays.

## Slide 4: Stroboscopic sampling
• No ADC samples a THz waveform directly; electronics stop at a few gigasamples per second.
• The laser fires an identical pulse every 10 ns (100 MHz repetition). The receiver gap conducts only during the picosecond it is illuminated, so each repetition samples the field at one delay.
• A motorized mirror sweeps the delay (1 mm of travel adds 6.67 ps of round trip). The readout measures a DC average current irec(Δt) proportional to the field at the gate instant; sweeping Δt reconstructs E(t).
Full TDS chain: one laser, a beam splitter, and a delay line (Lecture 2)
Notes: Electronics top out at Gsamples/s, not THz. Laser fires an identical pulse every 10 ns; receiver conducts only during the ps it's lit, one delay per repetition, set by a motorized mirror (1 mm = 6.67 ps). Readout: DC current proportional to field, so sweeping delay rebuilds the waveform.

## Slide 5: Noise: what the lecture predicts
• Average transmit power 1 mW at 100 MHz repetition: 10 pJ per pulse, roughly 20 W of peak power.
• With link efficiency near 1/10, the receiver current burst is about 0.14 A for half a picosecond.
• The duty cycle dilutes it: iave = ipeak · τrec/Trep ≈ 7 µA of DC current at the amplifier input.
• The readout amplifier noise is about 70 pA, so a single sweep already delivers a signal-to-noise ratio near 100 dB.
• Averaging N sweeps beats down white noise only:
SNRN = SNR1 + 10·log10N
Prediction for the lab
N = 50 → +17.0 dB
N = 100 → +20.0 dB
N = 1000 → +30.0 dB
Assignment 1 tests exactly this.
Notes: 1 mW at 100 MHz = 10 pJ/pulse, ~20 W peak. 10x link loss gives ~0.14 A burst for 0.5 ps, duty-cycle averaged to ~7 μA DC. Against 70 pA noise, ~100 dB in one sweep. Averaging: +10 log N, 17 dB at 50, 30 dB at 1000.

## Slide 6: The lab bench
Emitter (right), four TPX50 lenses, detector and LCA-S10 amplifier (left)
Slab at the pinhole focus
Emitter PCA on its carrier
• Both PCAs are fibre-coupled to the Menlo unit. Lenses 1 and 4 collimate at the antennas, lenses 2 and 3 form an intermediate focus where the samples go (carriers at 35 / 70 / 190 / 245 / 305 / 400 / 435 mm).
• Alignment by maximizing detector current lens by lens; the aligned pulse peaks at 2.89 a.u.
• Every scan: 100 ps delay window, 33.3 fs steps, Tukey window (α = 0.1) in ScanControl.
Notes: Emitter right, detector left, fibre-coupled to Menlo unit. Four TPX lenses: outer pair collimates at antennas, inner pair focuses at the pinhole. Aligned lens by lens, peak 2.89. Scan: 100 ps window, 33 fs steps.

## Slide 7: Averaging, measured against the prediction
• Dynamic range grows from 59.6 dB (N = 1) to 89.3 dB (N = 1000): gains of 16.9 / 19.3 / 29.7 dB against the predicted 17 / 20 / 30 dB. The white-noise law holds within 0.7 dB.
• The time trace saturates: past N = 100 the pre-pulse floor stays near 4·10-4 of the peak because a coherent baseline structure repeats in every sweep. Averaging only buys bandwidth where the noise is white.
Notes: Prediction holds: 60 dB unaveraged to 89 dB at 1000 averages. Measured 16.9/19.3/29.7 dB vs 17/20/30 predicted, within 0.7 dB. Past 100 averages, pre-pulse floor plateaus, coherent baseline repeats every sweep. What averaging really buys: bandwidth.

## Slide 8: Extracting permittivity and loss
• Differential measurement: a reference scan without the sample, a sample scan with it, and their ratio H(f) = Esam(f) / Eref(f). The N = 100 trace is the reference.
• Delay gives the permittivity: the slab replaces a thickness d of air, so the pulse arrives Δτ = (n−1)·d/c0 later and εr = n2.
• Loss comes from the spectrum: model the slab as a transmission-line section between air lines, then fit tan δ(f) until the modelled magnitude matches the measurement.
• The model contains the Fabry-Perot etalon of the slab faces. If the internal echo lands outside the 100 ps window (slabs 1 and 3), only the first transit is kept.
Air / slab / air transmission-line model from the processing notes
Notes: Every measurement is differential, scan without sample, scan with, take the ratio. Delay gives permittivity directly. Loss: model the slab as a transmission-line section, fit tan delta per frequency to match the transfer function, including the face etalon. Two thick slabs: echo falls outside the scan window, keep first transit only.

## Slide 9: Six slabs identified
Reference and slab pulses; the delay carries εr
Slab | d (mm) | εr | tan δ (1 THz) | Material
1 | 7.5 | 1.69 | 0.0010 | porous PTFE *
2 | 3.08 | 1.31 | 0.0026 | Gore-Tex
3 | 4.8 | 7.66 | 0.0131 | Ceramic
4 | 0.525 | 11.72 | 0.0138 | Silicon
5 | 0.525 | 11.66 | 0.0028 | Silicon
6 | 2.04 | 2.05 | 0.0082 | Teflon
• Five of six match the manual's reference table within 3.1%.
• Slab 3 arrives 28 ps late and heavily attenuated; the silicon wafers show their etalon echo 12 ps after the main pulse.
• * Slab 1 (εr = 1.69) sits between Gore-Tex and Teflon: consistent with porous PTFE of intermediate density.
Notes: Six slabs, delays 1.5-28 ps. Matched to the manual: two silicon wafers, a ceramic, Teflon, Gore-Tex, five of six within ~3%. Slab one is the outlier, likely porous PTFE, intermediate density.

## Slide 10: Transfer function fits
• The fit reproduces level and etalon ripple; on silicon the ripple period c0/(2nd) ≈ 83 GHz pins the optical thickness.
Notes: Ripple on the silicon wafers is the face etalon, period ~83 GHz, model reproduces it and pins the optical thickness independently of delay.

## Slide 11: The slab 2 thickness conflict
• The data file says 0.8 mm, the label on the holder says 3.08 mm. The delay alone cannot decide: both give a plausible material (εr = 2.44, HDPE-like, versus 1.31, Gore-Tex-like).
• Three independent tests reject 0.8 mm: the fit residual is three times worse, the required ±10% ripple at 120 GHz is absent, and the predicted internal echo at +8.3 ps (0.13 a.u.) does not exist in the residual.
• Conclusion: d = 3.08 mm and the slab is Gore-Tex. The file name recorded a different sample or a typo.
Notes: Slab 2's file says 0.8 mm, sample labelled 3.08. Delay alone can't decide. Three checks reject 0.8: fit residual 3x worse, expected ripple absent, predicted echo at 8.3 ps after the main pulse absent. So: 3.08 mm Gore-Tex, file name was wrong.

## Slide 12: Silicon wafers: same index, different loss
Extracted tan δ(f): silicon falls as 1/f, polymers rise
• Both wafers give n ≈ 3.42, within 0.3% of the literature value for high-resistivity silicon. The etalon ripple period confirms the permittivity independently.
• Their loss differs by an order of magnitude. Conduction loss follows tan δ = σ/(ω·ε0·εr), so constant conductivity appears as a 1/f slope. That is the measured behaviour below 1.5 THz.
• Converting the median conductivity of wafer 5 to resistivity gives 37 Ω·cm. The wafer is labelled 33 Ω·cm from a four-point-probe measurement: an independent end-to-end check of the loss extraction.
• Wafer 4 comes out at 9 Ω·cm, the more doped and lossier twin.
Notes: Two wafers: same index, 3.42, within 0.3% of literature, very different loss. Conduction loss: tan delta scales as 1/f, matches slope below 1.5 THz. Resistivity: wafer 5 = 37 Ω·cm vs labelled 33, within 12%. Wafer 4 ≈ 9, more doped.

## Slide 13: What this lab showed
• Stroboscopic sampling works as taught: a DC photocurrent measurement reconstructs the THz field with 33 fs steps.
• The averaging law holds: 29.7 dB measured against 30 dB predicted at N = 1000. What it cannot remove is coherent structure, and the time-domain floor shows exactly that.
• One delay reading plus a transmission-line fit identified five of six materials within 3.1% of the reference permittivities.
• The same physics settles a metadata conflict (slab 2) and reproduces a wafer's labelled resistivity within 12%.
Slab 2, marked t = 3.08 mm
Notes: Four things: stroboscopic sampling gives fs resolution from a DC measurement. Averaging follows 10 log N until coherent structure takes over. Delay plus transmission-line fit IDs materials to a few percent, and catches bad metadata along the way.

## Slide 14
Thank you
Time-Domain topic | Daniel Tyukov | EE4730 oral exam
Notes: Thank you. Happy to take questions.